FLEXA Digital Pirani Vacuum Gauge FGE-2GH

The UVS Model FLEXA series Vacuum Pirani Gauge Controller displays vacuum pressure as measured from a FG-04 sensor tube. The Model FGE-2GH unit covers the range from 0.001 m.Bar to 20 m.bar, and measures two channels.

Exponential order display from 2.0X101 to 1.0X10-3 mbar

The Advantage Of This Vacuum Gauge FLEXA Series Is The Sensor Tube Has The Built In Zero And Atm Adjustments, So That The Sensor Tube Can Be Used As A Transducer, And Each Sensor Is Factory Calibrated At Sensor Level.

With This 1:1 Replacement Of Sensor With The Display Unit Is Established, And There Is No Dependency On Vacuum Gauge Level Adjustment In Case Of Sensor Replacement.

TECHNICAL SPECIFICATIONS

| Number of sensor channels | TWO |
| --- | --- |
| Display Range | 1000 mBar to 0.001mBar. |
| Measuring Range | 20 mBar to 0.001mBar. |
| Units of Display | m.Bar |
| Vacuum Gauge Head | FG-04 sensor tube |
| Temp. Range | +2 to +50 deg. Celsius |
| VAC. Display | 4-digit bright red LED, 14 mm high |
| VAC and ATM Adjust | adjusted by trim pot at the signal conditioner PCB of the sensor. |
| Process control set points | NIL |
| Process Control Relays | NIL |
| Nonvolatile Memory | NIL |
| Operating Voltage | universal power supply, input voltages from 90 ~ 264VAC 50/60 Hz |
| Mounting | ¼ DIN 92 x 92 mm (3.62” x 3.62") |
| Environmental Considerations | not for use with explosive or corrosive gases |
| Overall dimensions | 92mm X 92mm X 110mm Depth |
| Optional outputs on request | RS485, 0 -10V Dc analog output (OPTIONAL) |
| Sensor cable | Standard RJ45 male plugs both ends- 2 mtrs long. Optional additional lengths available on request. |

Features at a Glance : ■ Measurement range 0.001 to 20 mbar ■ Two channel measurement ■ Space saving rugged design- 1/8 DIN console ■ Mounts in any orientation ■ brass body measuring cell with metal-sealed feed through ■ interchangeability of the sensor possible without disturbing the display unit.

Applications: ■ Controlling high vacuum systems ■ Fore vacuum pressure monitoring ■ Safety circuits in vacuum systems ■ General vacuum measurement & control in the fine and rough vacuum range.

Calibration:

Unique Vacuum Solutions offers a convenient recalibration service to help users of vacuum measurement pirani/ thermocouple/ full range gauges to maintain all of their vacuum gauges in Certified and peak condition, regardless of the original manufacturer. All makes of vacuum gauges can also be calibrated / repaired if required.

The calibration certificates are available only on request at additional cost. Pl contact before ordering.

UNIQUE VACUUM SOLUTIONS

#57, 8th cross, 1st B cross, Doddanna industrial area main road, vishwaneedam post, Near Peenya 2nd stage, Bangalore -560091

Ph: 080-28367059, 9886726920. +91-support@uniquevacuum.co.in, www.uniquevacuum.co.in