FLEXA Digital Pirani Vacuum Gauge  FGE-2GH-SP

It is two channel readout vacuum gauge with relay set point controller output facility.

The UVS Model FLEXA series Vacuum Pirani Gauge Controller displays vacuum pressure as measured from a FG-04 sensor tube. The Model FGE-2GH-SP displays vacuum measurements from 0.001mBar to 20 mbar, and has a built in potential free contact set point relay output, one for each channel.

The Advantage Of This Vacuum Gauge FLEXA Series Is The Sensor Tube Has The Built In Zero And Atm Adjustments, So That The Sensor Tube Can Be Used As A Transducer, And Each Sensor Is Factory Calibrated At Sensor Level.

With This 1:1 Replacement Of Sensor With The Display Unit Is Established, And There Is No Dependency On Vacuum Gauge Level Adjustment In Case Of Sensor Replacement.

| Sensor channels | TWO |
| --- | --- |
| Display Range | 1000 mBar to 0.001mBar. |
| Measuring Range | 20 mBar to 0.001mBar. |
| Measuring unit | m.Bar |
| Vacuum Gauge Head | FG-04 sensor tube |
| Temp. Range | +2 to +50 deg. Celsius |
| VAC. Display | 4-digit bright red LED, 14 mm high |
| VAC and ATM Adjust | adjusted by trim pot at the signal conditioner PCB of the sensor. |
| Process control set points | UPPER LEVEL & LOWER LEVEL set points for each channel. |
| Process Control Relays | two relay outputs, potential free contacts; independent available through the DB15 accessory connector on the rear panel |
| Nonvolatile Memory | Only for set points |
| Operating Voltage | universal power supply, input voltages from 90 ~ 264VAC 50/60 Hz |
| Mounting | ¼ DIN 92 x 92 mm (3.62” x 3.62") |
| Environmental Considerations | not for use with explosive or corrosive gases |
| Overall dimensions | 92mm X 92mm X 110mm Depth |
| Optional outputs on request | RS485, 0 -10V Dc analog output (optional) |
| Sensor cable | Standard RJ45 male plugs both ends- 2 mtrs long. Optional additional lengths available on request. |

Features at a Glance : ■ Measurement range 0.001 to 20 mbar ■ Space saving rugged design- 1/8 DIN console ■ brass body measuring cell with metal-sealed feed through ■ with universal power supply ■ interchangeability of the sensor possible without disturbing the display unit.

Applications: ■ Controlling high vacuum systems ■ Fore vacuum pressure monitoring ■ Safety circuits in vacuum systems ■ General vacuum measurement & control in the fine and rough vacuum range.

Calibration:

Unique Vacuum Solutions offers a convenient recalibration service to help users of vacuum measurement pirani/ thermocouple/ full range gauges to maintain all of their vacuum gauges in Certified and peak condition, regardless of the original manufacturer. All makes of vacuum gauges can also be calibrated / repaired if required.

The calibration certificates are available only on request at additional cost. Pl contact before ordering.

UNIQUE VACUUM SOLUTIONS

#57, 8th cross, 1st B cross, Doddanna industrial area main road, vishwaneedam post, Near Peenya 2nd stage, Bangalore -560091

Ph: 080-28367059, +91-9886726920. support@uniquevacuum.co.in, www.uniquevacuum.co.in