Pirani Vacuum Sensor FG-05A

The UVS Model FG-05A series Vacuum Pirani Gauge measurement of vacuum covers the range from 0.001 m.Bar to 20 m.bar.

The advantage of this vacuum gauge FLEXA series is the sensor tube has the built in zero and atm adjustments, so that the sensor tube can be used as a transducer, and each sensor is factory calibrated at sensor level.

With This 1:1 Replacement Of Sensor With The Display controller Unit Is Established, And There Is No Dependency On Vacuum Gauge Level Adjustment In Case Of Sensor Replacement.

Pirani Vacuum Sensor FG-05A

TECHNICAL SPECIFICATIONS :

| Vacuum Gauge Head | FG-05A Sensor Tube |
| --- | --- |
| Measuring Range | 20 Mbar To 0.001mbar. |
| Temp. Range | +2 To +50 Deg. Celsius |
| Vac And Atm Adjust | Adjusted By Trim Pot At The Signal Conditioner Pcb Of The Sensor. |
| Operating Voltage | 24V Dc |
| Mounting | Vertical |
| Linear Output | 1 Volt Per Decade |
| Sensor Cable Connection | 4 Pin Metal Sheel Connector. |
| Output Voltage | 0-10V Dc |
| Mass | g |
| Environmental Considerations | Not For Use With Explosive Or Corrosive Gases |
| Internal Volume | Cc |
| Accuracy | Typically +/- 15% At <20 Mbar |
| Maximum Over-Pressure | 2 Bar Absolute |
| Power Consumption | 1 W |
| Output Signal | 0 To 10 V d.c. Nominal |
| Indications | Green Led Status Indicator Shows Normal And Red Led Status Indicator Shows Fault / Fail Condition |

Applications: ■ Controlling high vacuum systems ■ Fore vacuum pressure monitoring ■ Safety circuits in vacuum systems ■ General vacuum measurement & control in the fine and rough vacuum range.

Calibration:

Unique Vacuum Solutions offers a convenient recalibration service to help users of vacuum measurement pirani/ thermocouple/ full range gauges to maintain all of their vacuum gauges in Certified and peak condition, regardless of the original manufacturer. All makes of vacuum gauges can also be calibrated / repaired if required.

The calibration certificates are available only on request at additional cost. Pl contact before ordering.

UNIQUE VACUUM SOLUTIONS

#57, 8th cross, 1st B cross, Doddanna industrial area main road, vishwaneedam post, Near Peenya 2nd stage, Bangalore -560091

Ph: 080-28367059, 9886726920. +91-support@uniquevacuum.co.in, www.uniquevacuum.co.in